# Pulse trains to percepts: The challenge of creating a perceptually intelligible world by electrically stimulating visual cortex

Media summary

For many blind individuals, one promising approach for restoring vision is direct electrical stimulation of visual cortex, using either surface or penetrating electrodes. One visual cortical implant based on surface electrodes (Orion, Second Sight Medical Implants) has already received conditional FDA approval, with others, based on penetrating, electrodes scheduled to start clinical trials within a few years. However, as yet we have little insight into the perceptual experience likely to be produced by these implants – what will the world look like to patients? Here we describe a simple model describing how interactions between the implant electronics and the underlying neurophysiology are likely to influence perceptual outcomes in visual cortical prosthesis patients.

Ione Fine*, Department of Psychology, University of Washington, ionefine@uw.edu

Michael Beyeler, Department of Psychology, University of Washington, ionefine@uw.edu

Ariel Rokem, XXX, University of Washington, gboynton@uw.edu

Geoffrey Boynton, Department of Psychology, University of Washington, gboynton@uw.edu

* Corresponding author

# Summary

Damage to the optic nerve or lateral geniculate nucleus is responsible for most cases of non-treatable blindness. For these individuals, one promising approach for restoring vision is direct electrical stimulation of visual cortex, using either surface or penetrating electrodes. One visual cortical implant based on surface electrodes (Orion, Second Sight Medical Implants) has already received conditional FDA approval, with others, based on penetrating, electrodes scheduled to start clinical trials within a few years. Interactions between the implant electronics and the underlying neurophysiology are likely to be an even more important factor for cortical implants, yet their influence has not been fully considered. Here we present a simple stimulation based on well-established models of V1 cortical physiology demonstrating how the underlying neurophysiology of V1 is likely to likely to influence perceptual outcomes in visual cortical prosthesis patients.

# Keywords

neural prosthetic

vision restoration

optogenetic

cortical stimulation

glaucoma

blindness

vision loss

# Introduction

Non-treatable vision loss affects over 10 million individuals worldwide {Retnet, 2014 #6973}. As a consequence, there is considerable interest in developing technologies to restore visual function. Although a variety of approaches, such as optogenetics, small molecule photoswitches and retinal prosthetics are currently being developed to restore vision in individuals suffering from vision loss due to retinal disease {Fine, 2015 #11638} the large majority individuals with non-treatable severe blindness suffer from diseases, such as glaucoma, that cause damage distal to the retina. This has motivated interest in developing electrical cortical implants, that could directly stimulate early visual areas {Dobelle, 1979 #7328}. One visual cortical implant based on surface electrodes (Orion, Second Sight Medical Implants) is already being implanted in human subjects, and at least one other technology is expected to reach human trial stage by 2022.

In the case of retinal prosthetics, once devices were implanted in patients it was quickly discovered that there were substantial perceptual distortions due to interactions between the implant electronics and the underlying neurophysiology. These interactions are likely to be an even more important factor for cortical implants, yet their influence has not been fully considered.

Cortical prosthesis development is currently implicitly driven by a ‘linear scoreboard’ model in which it is assumed that the size of an electrically elicited phosphene can be predicted by the projection of the electrical field profile on the cortical surface into visual space co-ordinates. According to this model, small electrodes will produce small discrete percepts.

Here we present a simple stimulation, based on well-established models of V1 cortical physiology that relies on simple well-established facts about orientation tuning, ocular dominance and receptive field sizes in V1. We find that the predicted effects of cortical electrical stimulation differ dramatically from ‘linear scoreboard’ model.

For example, even in the fovea, the majority of neurons have receptive fields are ~1° of visual angle. Without a way of selectively stimulating neurons with small receptive fields or particular phase preferences, neuronal receptive field size rather than electrode size will limit perceptual resolution unless electrodes are separated by many millimeters of cortex.  Indeed, smaller (<0.5mm) electrodes will selectively stimulate orientation and (depending on electrode depth) ocular dominance columns. As a result, this is likely to result in highly confusing perceptual experiences, analogous to presenting binocularly rivalrous stimuli (different stimuli in each eye).

Results

# Further concerns: distortions due to retinal cell diversity

Signals from mammalian cones flow to a variety of morphologically distinct (at least three cell types [33]) “OFF-centre” and “ON-centre” bipolar cells. These pathways pass in turn to ON- and OFF-centre ganglion cell pathways that each contain of 10 or more cell types [31, 32]. This diversity within ganglion cells is in striking contrast to the auditory nerve fibres stimulated by cochlear implants, whose responses show far less diversity in functional response [e.g. 34, 35]. Each type of bipolar [33] and ganglion cell has strikingly different functional properties and carries distinct information about the visual input [32]. In the case of ganglion cells, these cell types form orderly mosaics [36-39], such that a given location in the visual field is represented by a wide range of cell types whose responses to a given stimulus vary dramatically (indeed stimuli that induce firing in ON-centre pathways tend to suppress OFF-centre pathways and vice versa).

Although the functional properties and cortical projections of many (though not all) of these cell types have been described, there is still much to be learned about how the information from these different retinal representations are combined at later stages of processing [see 40, for a review]. One reason for this is despite rapid progress [41], it remains extremely difficult to selectively stimulate or block particular cell types or ON- or OFF-centre cells, especially in primates. However it is clear that these representations are not independent [42, 43] and are often (likely due to the rectified nature of neural responses) complementary. This diversity of representation within bipolar and ganglion cells has significant implications for sight recovery technologies. Although optogenetics and small molecule photoswitches can be targeted with varying degrees of cell specificity (e.g. to cones [44], ganglion cells [4, 5, 7, 12-14], bipolar ON cells [7, 13, 15], and even ganglion cell dendrites and soma [45]), none of the technologies currently under development have the ability to match the normal pattern of responses across the full mosaic of cell types. To achieve this would require multiple compounds, each capable of being independently stimulated (e.g. each narrowly tuned to a different wavelength of light) and each capable of selectively targeting a particular cell type (or a small subset of cell types with similar response properties). What then is the likely perceptual consequence of targeting a subset of the full cell mosaic?

## The effect of stimulating either the ON- or the OFF- pathway

First we consider the potential perceptual consequences of targeting ON- or OFF-centre pathways in isolation. Figure 1 Panel A shows two example grayscale images, and Panels B and C show simple approximations of the representations thought to be carried by ON- and OFF-pathways. While perceptually interpretable, neither look much like the original image, and both carry impoverished information content compared to the band-pass image shown in Panel D.

Figure 1. A simulation of the perceptual information carried by ON- and OFF-centre pathways. (Column A) Two example images: a white-on-black eye-chart and a frame of a natural scene movie of a child scooting. Both subtend 24 degrees of visual angle. (Column B) These two images were convolved by a bank of difference of Gaussian filters followed by rectification to provide a rough approximation of the information carried by ON-centre pathways. (Column C) Images filtered by an inverted version of the same bank of Gaussian filters followed by rectification provides a rough approximation of the information carried by OFF-centre pathways. (Column D) The difference of the images of Panels B and C restore a band-pass version of the original image. (Column E) The sum of the images Panels B and C, represents an approximation of one potential perceptual outcome of simultaneously stimulating both ON-centre and OFF-centre pathways. Movie versions are shown for Panels A-E in Supplementary Materials.

However this simulation underestimates the complexity of stimulating ON- but not OFF-pathways (or vice versa), because it doesn’t represent the fact that these pathways are complementary to each other. An absence of response in a given pathway is not at all the same thing as an absence of information from that pathway. To take a simple example, if one somehow blocked transmission in all leftward selective motion sensitive cells, then it would not be the case that one simply wouldn’t be able to see leftward motion (the analogy to the simulation above) – rather, due to the nature of motion opponency it is likely that everything would be perceived as moving rightward. Interestingly, while it is possible to visualize or simulate the perceptual results of blocking transmission of leftward tuned cells, it is impossible to visualize or stimulate the perceptual result of stimulating only ON-centre or OFF-centre pathways because there is no visual stimulus that elicits such a retinal response – responses to visual stimuli in ON-cells are always accompanied by suppression in OFF-centre cells with receptive fields in nearby locations. Indeed, one possibility is that this type of stimulation will create a ‘rivalrous’ stimulus (as when different stimuli are presented to each eye [46]) in which one pathway signals a blank field and the other signals some sort of visual stimulus, as discussed more fully below.

## The effects of unselective stimulation within ON- and OFF- pathways

Even if just OFF- or ON- pathways are selectively stimulated, any stimulation protocol will require a compromise across cell types whose functional response properties differ in terms of their receptive field size, transience and chromatic tuning. Thus, a stimulation protocol that matches the average or the modal rate of cells within the ON- or OFF- pathway, will provide a remarkably poor description of the firing of any individual cell.

Thus for optogenetics and small cell photoswitches the choice of targets may prove a ‘devil’s choice’. If cells with too broad a range of functional properties are stimulated then the information carried by spikes may become temporally garbled. If too narrow a collection of cells is stimulated then the perceptual representation may prove too highly impoverished.

## The effect of stimulating both ON- and OFF- pathways

Current electrical prosthetics stimulate a wide range of cells (ganglion, bipolar and amacrine cells, also ganglion cell axon fibres, as discussed in the next section). Moreover, the spiking pattern elicited by electrical stimulation is similar across all these cells, regardless of whether the cell is ON or OFF centre and regardless of cell type (parasol, midget etc.). This pattern of simultaneous firing is strikingly different from the normal responses of the retina. Indeed, the simultaneous firing of both ON- and OFF- ganglion cells representing the same location is something that never happens naturally.

An example of the potential perceptual consequences of simultaneously activating ON- and OFF- pathways can be seen in Figure 1E. As described above, Panels B and C show simple approximations of the representations thought to be carried by ON- and OFF- pathways. When combined appropriately the two representations produce a band-pass version of the original image (Panel D). However, when combined with a sign-flip, as would occur if both ON- and OFF- cells were identically stimulated (with the additional presumption that equal stimulation in ON- and OFF- pathways ‘cancels each other out’) the resulting perceptual image is a ‘cartoon sketch’ of outlines (Panel E). Alternatively, as mentioned above, this pattern of stimulation, in which contradictory information is carried by ON- and OFF- channels, may create a ‘rivalrous’ stimulus (similar to the binocular rivalry that occurs when different images are presented to each eye [46]) in which the percept might either be dominated by either the ON- or the OFF-centre pathway, or might dynamically shift between the representation carried by each pathway.

# Axonal stimulationRetinal prosthetics can either be implanted subretinally next to the choroid, in the space of the missing or ailing photoreceptors, or epiretinally, between the ganglion cells and the vitreous humour. Every ganglion cell has an axon that traverses the retinal surface en route to the optic nerve, between the ganglion cell bodies and the vitreous humour, so in the case of epiretinal implantation, the axon fibres lie between the electrodes and cell bodies. Both animal models [47], modelling [48] and human data [49] suggest that epiretinal devices produce significant axonal stimulation. Stimulating a ganglion cell axon is likely to lead to the percept of a stimulus at the location associated with that axon’s cell body, which could be several degrees away. Thus an electrode that stimulates underlying axonal fibres would be expected to produce the percept of ‘comets’ whose head lies at the intended location of stimulation and whose tail lies along the axonal fibre pathway in the direction leading away from the optic disc, as illustrated in Figure 2A. As a consequence, if an electrode stimulates axons of ganglion cells with distant cell bodies, then the percepts elicited by electrical stimulation of those axons will be elongated in shape and poorly localized.

As described previously [19], Argus I and II (epiretinal arrays) patients typically report that phosphenes appear light grey, white or yellow with the shape varying between round to highly elongated ellipses. As the stimulation amplitude or frequency increases, subjects tend to report phosphenes as brighter with sharper contours. Suggestive of axonal stimulation, nearly all phosphenes appear as elongated ellipses (∼93%), with their minor axis length being less than 50% of the major axis length.

To simulate the perceptual effects of axonal stimulation, we used a previously validated model of retinal axonal pathways based on that of Nanduri et al., [50]. Four Argus I and Argus II retinal prosthesis subjects were asked to trace the shape of induced phosphenes on a touch screen. These perceptual data were fit using a model that assumed that the activation sensitivity of a passing axon fibre decays exponentially with the distance (x) between the axon’s initial segment and the electrode, such that s=e-x/λ. A constant of λ = ∞ implies that the entire axon fibre is equally sensitive to electrical stimulation, λ = 0 means only the initial segment is sensitive to stimulation. Nanduri et al. found that patient percepts were best predicted using values of λ ranging between 1.0 and 3.4 degrees of visual angle, indicating that while electrical stimulation is not confined to the axon initial segment, it did fall off rapidly as a function of distance.

As shown in Figure 2, this model was able to successfully predict the length and orientation of phosphenes based on the orientation and length of the underlying axon fibres. In addition, the model could successfully predict whether stimulation of two electrodes resulted in a single or double percept, based on whether or not the two electrodes fell along shared fibre pathways.

Figure 2. An example of how a model of retinal axonal pathways can predict patient percepts. (A) Schematic of why axonal stimulation results in axon comets. Each image of the retina is flipped so that the upper region of the retina represents the upper visual field. The modelled axon trajectories (green lines) are based on a computational model of axon fibre trajectories developed using traced nerve fibre bundle trajectories extracted from fundus photographs of 55 human subjects. [51]. The red circles represent ganglion cell bodies whose axon fibres pass underneath the electrode on their way to the optic nerve, yellow shading represents a perceptual ‘axon-comet’. The dotted red box outlines the retinal regions shown in Panels B-D. (B, C) A subject implanted with the Argus 1 prosthesis (this array contained interleaved 250 and 500m electrodes) was simultaneously stimulated on two different pairs of electrodes. Each pair of stimulated electrodes are shown outlined in red. Predicted percepts generated using the model described above are shown in white. (D, E) Subject drawings (averaged over 5 trials) of the percepts induced by these stimulation patterns.

While the increased distance of the stimulating electrodes means that axonal stimulation may be less of a concern for subretinal devices [52, 53], there are indications that in a subset of patients a certain amount of axonal stimulation occurs, especially at suprathreshold levels of stimulation While most patients implanted with subretinal electrodes report percepts as round spots of light with a yellowish appearance, some patients report percepts that include arcs, short lines or semicircles [52, 53].

In Figure 3 we use this model to show examples of the perceptual distortions that might be produced by axonal stimulation with more naturalistic images, using a 101 x 101 square electrode array subtending 12 degrees of visual angle. Here we show simulations based on values of λ = 0.5, 1 and 2 degrees of visual angle. Panels A and B show the images projected onto the retina, overlaid on the simulated axon fibre pathways. Panels C and D show the effect of such axonal ‘comets’ on specific images. Since distortions follow the path of the axon fibre pathways, the ‘comet-trails’ predicted by axonal stimulation vary in angle and length across the retina In the simulation of a contrast-reversed (white letters on a black background) eye-chart using λ = 2 (based on [49]) only the top E (which subtends just less than 3 degrees of visual angle) is clearly identifiable, consistent with Snellen acuity of 20/600, similar to the best acuity reported to date (just under 20/600) for human prosthetic users [10, 17].

Figure 3. Simulations of perceptual distortions as a result of axonal stimulation. (A, B) Two images (the same as in Figure 1) are overlaid on the retinal surface. The position of the electrode array (subtending 12 degrees) is shown by a red dotted box. (C, D) The predicted effect of axonal stimulation on the two images of Figure 1B and D for λ = 0.5, 1 and 2. Movie versions are shown in Supplementary Materials.

The simulation of the natural image is still more troublesome: for λ=2 even the main figure in the image cannot clearly be delineated. The critical difference seems to be that in the contrast-reversed letter chart only a small subset of the electrodes are stimulated, and the electrodes representing the black background remain entirely unstimulated. In contrast most electrodes are partially activated in the natural scene, resulting in ‘comet-trails’ smearing the image much more severely. This is a particular concern given that the vast majority of the paradigms used to evaluate both acuity and ‘real world performance’ in retinal implant users have relied on discrete white objects on a black background [e.g. 8, 17, 54], Figure 4.

Figure 4. Example of one test within a standardized ‘Activities of daily living’ test proposed by [54, reprinted with their kind permission]. The subject is asked to identify, describe and localize objects while sitting at a table.

One way of reducing the perceptual consequences of comet trails might to carry out initial image processing to enhance contrast and edges (e.g. ‘zero-crossing’ algorithms), and minimize the number of stimulated electrodes. Alternatively it is possible that these perceptual effects can be reduced by incorporating them into the encoding model, such that the visual region encoded by an electrode is represented by its ‘comet’ rather than by the region of visual space corresponding to position of the electrode on the retinal surface. However, even if the phosphenes elicited by each electrode were accurately represented, creating small discrete phosphenes (such as spots, or lines) would remain impossible, even with a very high resolution array. For example, in Figure 2C and E it is highly ambiguous which and how many electrodes along the fibre pathway are being stimulated.

A third possibility is developing more complex encoding models that include anodic (suppressive) stimulation on electrodes on the same axon pathway but distal to the optic nerve. Thus, to create a punctate phosphene under the electrode of Figure 2A one would anodically stimulate electrodes lying above the ganglion cells represented by red circles. Note that while this technique is related to ‘current shaping’ (in which a combination of anodic and cathodic stimulation on electrodes with overlapping current fields are used to ‘shape’ the current field [55-58]) it is conceptually different because the interactions between electrodes are neural rather than within the current field. The electrodes used to suppress axonal fibre stimulation need not have overlapping current fields with the stimulating electrode, but rather must lie along the same axon pathway.

# Conclusions

As a final example, we show three simulations. All are once again based on a 101 x 101 array of electrodes subtending 12 degrees. Figure 6A represents a basic score-board model in which the brightness of the apparent percept is linearly related to the strength of current on the retina (based on [86]). Panels B and C represent ‘neuro-perceptual’ simulations of the effects of electrical and small molecule photoswitch stimulation respectively. Each simulation is strikingly different, and it is immediately apparent that array resolution may be less critical than the ability to create perceptually comprehensible percepts.

Figure 6. Three example models of the potential perceptual experience of sight recovery. All images subtend 12 degrees of visual angle. (A) Scoreboard model. The position of the array is again shown with red dotted lines. The luminance of the apparent percept is linearly related to the strength of current on the retina. (B) Simulation of electrical stimulation. This particular simulation is based on the model of simultaneously stimulating ON- and OFF pathways as described in Figure 1E, followed by the model of axonal stimulation as described in Figure 3 (C) Simulation of small molecule photoswitch stimulation. This simulation is based on the model of simulating ON-centre pathways in isolation as described in Figure 1B, followed by the effects of sluggish temporal dynamics as described in Figure 5. Movie versions are shown in Supplementary Materials.

Thus, the next decade is likely to see fascinating interdisciplinary research examining the interplay between sight recovery technologies, the underlying neurophysiology and the perceptual capacities of individuals. As second generation sight recovery technologies begin to develop encoding schemes to maximize the perceptual intelligibility of the world, we believe ‘neuro-perceptual’ models of the effects of artificial stimulation will play an increasingly important role. The simulations of this paper should not be considered as having the status of models, or of providing genuine predictions of what might be expected of any given sight recovery technology. Rather they are an attempt to demonstrate the useful information that more fully developed and validated models might provide.

# Acknowledgments

Many thanks to EJ Chichilnisky, Sheila Nirenberg and Bruce Cummings for helpful discussions and comments on the manuscript. Also many thanks to Eberhardt Zrenner, and John Flannery for generous permission to reprint a figure and data from their papers. Supported by NIH EY-12925 and EY-014645.

# Competing Interests

We have no competing interests.

# Funding

This work was supported by National Institutes of Health Grants No. R01EY-014645 (I.F.) and R01EY-12925 (G.M.B.).

# Authors' contributions

IF, MB, AR and GMB amicably participated in model development, figure creation and writing the manuscript. All authors gave final approval for publication.

# References

[1] Retnet. 2014 Genes and Mapped Loci Causing Retinal Diseases.

[2] Bamann, C., Nagel, G. & Bamberg, E. 2010 Microbial rhodopsins in the spotlight. Current opinion in neurobiology 20, 610-616. (doi:10.1016/j.conb.2010.07.003).

[3] Bi, A., Cui, J., Ma, Y.P., Olshevskaya, E., Pu, M., Dizhoor, A.M. & Pan, Z.H. 2006 Ectopic expression of a microbial-type rhodopsin restores visual responses in mice with photoreceptor degeneration. Neuron 50, 23-33. (doi:10.1016/j.neuron.2006.02.026).

[4] Busskamp, V., Picaud, S., Sahel, J.A. & Roska, B. 2012 Optogenetic therapy for retinitis pigmentosa. Gene Ther 19, 169-175. (doi:10.1038/gt.2011.155).

[5] Tochitsky, I., Polosukhina, A., Degtyar, V.E., Gallerani, N., Smith, C.M., Friedman, A., Van Gelder, R.N., Trauner, D., Kaufer, D. & Kramer, R.H. 2014 Restoring visual function to blind mice with a photoswitch that exploits electrophysiological remodeling of retinal ganglion cells. Neuron 81, 800-813. (doi:10.1016/j.neuron.2014.01.003).

[6] Polosukhina, A., Litt, J., Tochitsky, I., Nemargut, J., Sychev, Y., De Kouchkovsky, I., Huang, T., Borges, K., Trauner, D., Van Gelder, R.N., et al. 2012 Photochemical restoration of visual responses in blind mice. Neuron 75, 271-282. (doi:10.1016/j.neuron.2012.05.022).

[7] Gaub, B.M., Berry, M.H., Holt, A.E., Reiner, A., Kienzler, M.A., Dolgova, N., Nikonov, S., Aguirre, G.D., Beltran, W.A., Flannery, J.G., et al. 2014 Restoration of visual function by expression of a light-gated mammalian ion channel in retinal ganglion cells or ON-bipolar cells. Proceedings of the National Academy of Sciences 111, E5574-E5583. (doi:10.1073/pnas.1414162111).

[8] Humayun, M.S., Dorn, J.D., da Cruz, L., Dagnelie, G., Sahel, J.A., Stanga, P.E., Cideciyan, A.V., Duncan, J.L., Eliott, D., Filley, E., et al. 2012 Interim results from the international trial of Second Sight's visual prosthesis. Ophthalmology 119, 779-788.

[9] Klauke, S., Goertz, M., Rein, S., Hoehl, D., Thomas, U., Eckhorn, R., Bremmer, F. & Wachtler, T. 2011 Stimulation with a wireless intraocular epiretinal implant elicits visual percepts in blind humans. Invest Ophthalmol Vis Sci 52, 449-455. (doi:10.1167/iovs.09-4410).

[10] Stingl, K., Bartz-Schmidt, K.U., Besch, D., Braun, A., Bruckmann, A., Gekeler, F., Greppmaier, U., Hipp, S., Hörtdörfer, G., Kernstock, C., et al. 2013 Artificial vision with wirelessly powered subretinal electronic implant alpha-IMS.

[11] Palanker, D., Vankov, A., Huie, P. & Baccus, S. 2005 Design of a high-resolution optoelectronic retinal prosthesis. Journal of Neural Engineering 2, S105.

[12] Mutter, M., Swietek, N. & Munch, T.A. 2014 Salvaging ruins: reverting blind retinas into functional visual sensors. Methods in molecular biology 1148, 149-160. (doi:10.1007/978-1-4939-0470-9_10).

[13] Lagali, P.S., Balya, D., Awatramani, G.B., Munch, T.A., Kim, D.S., Busskamp, V., Cepko, C.L. & Roska, B. 2008 Light-activated channels targeted to ON bipolar cells restore visual function in retinal degeneration. Nat Neurosci 11, 667-675.

[14] Zhang, Y., Ivanova, E., Bi, A. & Pan, Z.H. 2009 Ectopic expression of multiple microbial rhodopsins restores ON and OFF light responses in retinas with photoreceptor degeneration. The Journal of neuroscience : the official journal of the Society for Neuroscience 29, 9186-9196. (doi:10.1523/JNEUROSCI.0184-09.2009).

[15] Doroudchi, M.M., Greenberg, K.P., Liu, J., Silka, K.A., Boyden, E.S., Lockridge, J.A., Arman, A.C., Janani, R., Boye, S.E., Boye, S.L., et al. 2011 Virally delivered channelrhodopsin-2 safely and effectively restores visual function in multiple mouse models of blindness. Molecular therapy : the journal of the American Society of Gene Therapy 19, 1220-1229. (doi:10.1038/mt.2011.69).

[16] Caporale, N., Kolstad, K.D., Lee, T., Tochitsky, I., Dalkara, D., Trauner, D., Kramer, R., Dan, Y., Isacoff, E.Y. & Flannery, J.G. 2011 LiGluR restores visual responses in rodent models of inherited blindness. Molecular therapy : the journal of the American Society of Gene Therapy 19, 1212-1219.

[17] da Cruz, L., Coley, B.F., Dorn, J., Merlini, F., Filley, E., Christopher, P., Chen, F.K., Wuyyuru, V., Sahel, J., Stanga, P., et al. 2013 The Argus II epiretinal prosthesis system allows letter and word reading and long-term function in patients with profound vision loss. The British journal of ophthalmology 97, 632-636. (doi:10.1136/bjophthalmol-2012-301525).

[18] Stingl, K., Bartz-Schmidt, K.U., Besch, D., Gekeler, F., Greppmaier, U., Hortdorfer, G., Koitschev, A., Peters, T., Sachs, H., Wilhelm, B., et al. 2012 [What can blind patients see in daily life with the subretinal Alpha IMS implant? Current overview from the clinical trial in Tubingen]. Der Ophthalmologe : Zeitschrift der Deutschen Ophthalmologischen Gesellschaft 109, 136-141. (doi:10.1007/s00347-011-2479-6).

[19] Nanduri, D., Fine, I., Horsager, A., Boynton, G.M., Humayun, M.S., Greenberg, R.J. & Weiland, J.D. 2012 Frequency and amplitude modulation have different effects on the percepts elicited by retinal stimulation. Invest Ophthalmol Vis Sci 53, 205-214. (doi:10.1167/iovs.11-8401).

[20] Horsager, A., Greenwald, S.H., Weiland, J.D., Humayun, M.S., Greenberg, R.J., McMahon, M.J., Boynton, G.M. & Fine, I. 2009 Predicting visual sensitivity in retinal prosthesis patients. Invest Ophthalmol Vis Sci 50, 1483-1491. (doi:10.1167/iovs.08-2595).

[21] Horsager, A., Boynton, G.M., Greenberg, R.J. & Fine, I. 2011 Temporal interactions during paired-electrode stimulation in two retinal prosthesis subjects. Invest Ophthalmol Vis Sci 52, 549-557. (doi:10.1167/iovs.10-5282).

[22] Horsager, A., Greenberg, R.J. & Fine, I. 2010 Spatiotemporal interactions in retinal prosthesis subjects. Invest Ophthalmol Vis Sci 51, 1223-1233. (doi:10.1167/iovs.09-3746).

[23] Greenwald, S.H., Horsager, A., Humayun, M.S., Greenberg, R.J., McMahon, M.J. & Fine, I. 2009 Brightness as a function of current amplitude in human retinal electrical stimulation. Invest Ophthalmol Vis Sci 50, 5017-5025. (doi:10.1167/iovs.08-2897).

[24] de Balthasar, C., Patel, S., Roy, A., Freda, R., Greenwald, S., Horsager, A., Mahadevappa, M., Yanai, D., McMahon, M.J., Humayun, M.S., et al. 2008 Factors affecting perceptual thresholds in epiretinal prostheses. Invest Ophthalmol Vis Sci 49, 2303-2314. (doi:10.1167/iovs.07-0696).

[25] Nirenberg, S. & Pandarinath, C. 2012 Retinal prosthetic strategy with the capacity to restore normal vision. Proc Natl Acad Sci U S A 109, 15012-15017. (doi:10.1073/pnas.1207035109).

[26] Rubinstein, J.T. 2004 How cochlear implants encode speech. Curr Opin Otolaryngol Head Neck Surg 12, 444-448.

[27] Shannon, R.V. 2012 Advances in Auditory Prostheses. Current opinion in neurology 25, 61-66. (doi:10.1097/WCO.0b013e32834ef878).

[28] Pillow, J.W., Shlens, J., Paninski, L., Sher, A., Litke, A.M., Chichilnisky, E.J. & Simoncelli, E.P. 2008 Spatio-temporal correlations and visual signalling in a complete neuronal population. Nature 454, 995-999. (doi:10.1038/nature07140).

[29] Pillow, J.W., Paninski, L., Uzzell, V.J., Simoncelli, E.P. & Chichilnisky, E.J. 2005 Prediction and decoding of retinal ganglion cell responses with a probabilistic spiking model. J Neurosci 25, 11003-11013.

[30] Keat, J., Reinagel, P., Reid, R.C. & Meister, M. 2001 Predicting every spike: a model for the responses of visual neurons. Neuron 30, 803-817.

[31] Rodieck, R.W. 1998 The first steps in seeing. Sunderland, Mass., Sinauer Associates.

[32] Dacey, D. 2004 20 Origins of Perception: Retinal Ganglion Cell Diversity and the Creation of Parallel Visual Pathways. In The Cognitive Neurosciences Iii (ed. M.S. Gazzaniga), p. 281, MIT Press.

[33] DeVries, S.H. 2000 Bipolar Cells Use Kainate and AMPA Receptors to Filter Visual Information into Separate Channels. Neuron 28, 847-856. (doi:10.1016/s0896-6273(00)00158-6).

[34] Winter, I.M., Robertson, D. & Yates, G.K. 1990 Diversity of characteristic frequency rate-intensity functions in guinea pig auditory nerve fibres. Hear Res 45, 191-202.

[35] Heil, P., Neubauer, H. & Irvine, D.R.F. 2011 An improved model for the rate-level functions of auditory-nerve fibers. Journal of Neuroscience 31, 15424-15437.

[36] Field, G.D., Gauthier, J.L., Sher, A., Greschner, M., Machado, T.A., Jepson, L.H., Shlens, J., Gunning, D.E., Mathieson, K., Dabrowski, W., et al. 2010 Functional connectivity in the retina at the resolution of photoreceptors. Nature 467, 673-677.

[37] Frechette, E.S., Sher, A., Grivich, M.I., Petrusca, D., Litke, A.M. & Chichilnisky, E.J. 2005 Fidelity of the ensemble code for visual motion in primate retina. J Neurophysiol 94, 119-135.

[38] Devries, S.H. & Baylor, D.A. 1997 Mosaic arrangement of ganglion cell receptive fields in rabbit retina. J Neurophysiol 78, 2048-2060.

[39] Wassle, H., Peichl, L. & Boycott, B.B. 1981 Dendritic territories of cat retinal ganglion cells. Nature 292, 344-345.

[40] Callaway, E.M. 2005 Structure and function of parallel pathways in the primate early visual system. The Journal of physiology 566, 13-19. (doi:10.1113/jphysiol.2005.088047).

[41] Berdyyeva, T.K. & Reynolds, J.H. 2009 The dawning of primate optogenetics. Neuron 62, 159-160. (doi:10.1016/j.neuron.2009.04.011).

[42] Mastronarde, D.N. 1983 Interactions between ganglion cells in cat retina. J Neurophysiol 49, 350-365.

[43] Shlens, J., Field, G.D., Gauthier, J.L., Grivich, M.I., Petrusca, D., Sher, A., Litke, A.M. & Chichilnisky, E.J. 2006 The structure of multi-neuron firing patterns in primate retina. J Neurosci 26, 8254-8266. (doi:10.1523/JNEUROSCI.1282-06.2006).

[44] Busskamp, V., Duebel, J., Balya, D., Fradot, M., Viney, T.J., Siegert, S., Groner, A.C., Cabuy, E., Forster, V., Seeliger, M., et al. 2010 Genetic reactivation of cone photoreceptors restores visual responses in retinitis pigmentosa. Science 329, 413-417.

[45] Greenberg, K.P., Pham, A. & Werblin, F.S. 2011 Differential targeting of optical neuromodulators to ganglion cell soma and dendrites allows dynamic control of center-surround antagonism. Neuron 69, 713-720. (doi:10.1016/j.neuron.2011.01.024).

[46] Wheatstone, C. 1838 Contributions to the Physiology of Vision. Part the First. On Some Remarkable, and Hitherto Unobserved, Phenomena of Binocular Vision. Philosophical Transactions of the Royal Society of London 128, 371-394. (doi:10.1098/rstl.1838.0019).

[47] Jensen, R.J., Ziv, O.R. & Rizzo, J.F., 3rd. 2005 Thresholds for activation of rabbit retinal ganglion cells with relatively large, extracellular microelectrodes. Invest Ophthalmol Vis Sci 46, 1486-1496.

[48] Tsai, D., Chen, S., Protti, D.A., Morley, J.W., Suaning, G.J. & Lovell, N.H. 2012 Responses of retinal ganglion cells to extracellular electrical stimulation, from single cell to population: model-based analysis. PLoS One 7, e53357. (doi:10.1371/journal.pone.0053357).

[49] Nanduri, D., Fine, I., Greenberg, R., Horsager, A., Boynton, G.M. & Weiland, J. 2011 Predicting the percepts of electrical stimulation in retinal prosthesis subjects. In Cosyne (

[50] Nanduri, D., Fine, I., Greenberg, R.J., Horsager, A., Boynton, G.M. & Weiland, J.D. 2011 Predicting the percepts of electrical stimulation in retinal prosthesis subjects. In Computational and Systems Neuroscience Meeting (

[51] Jansonius, N.M., Nevalainen, J., Selig, B., Zangwill, L.M., Sample, P.A., Budde, W.M., Jonas, J.B., Lagreze, W.A., Airaksinen, P.J., Vonthein, R., et al. 2009 A mathematical description of nerve fiber bundle trajectories and their variability in the human retina. Vision Res 49, 2157-2163. (doi:10.1016/j.visres.2009.04.029).

[52] Wilke, R., Gabel, V.P., Sachs, H., Bartz Schmidt, K.U., Gekeler, F., Besch, D., Szurman, P., Stett, A., Wilhelm, B., Peters, T., et al. 2011 Spatial resolution and perception of patterns mediated by a subretinal 16-electrode array in patients blinded by hereditary retinal dystrophies. Invest Ophthalmol Vis Sci 52, 5995-6003. (doi:10.1167/iovs.10-6946).

[53] Tsai, D., Morley, J.W., Suaning, G.J. & Lovell, N.H. 2009 Direct activation and temporal response properties of rabbit retinal ganglion cells following subretinal stimulation. J Neurophysiol 102, 2982-2993. (doi:10.1152/jn.00545.2009).

[54] Stingl, K., Bach, M., Bartz-Schmidt, K.U., Braun, A., Bruckmann, A., Gekeler, F., Greppmaier, U., Hortdorfer, G., Kusnyerik, A., Peters, T., et al. 2013 Safety and efficacy of subretinal visual implants in humans: methodological aspects. Clinical & experimental optometry : journal of the Australian Optometrical Association 96, 4-13. (doi:10.1111/j.1444-0938.2012.00816.x).

[55] Donaldson, G.S., Kreft, H.A. & Litvak, L. 2005 Place-pitch discrimination of single- versus dual-electrode stimuli by cochlear implant users. J Acoust Soc Am 118, 623-626.

[56] Hughes, M.L. & Goulson, A.M. 2011 Electrically evoked compound action potential measures for virtual channels versus physical electrodes. Ear Hear 32, 323-330.

[57] Bonham, B.H. & Litvak, L.M. 2008 Current focusing and steering: modeling, physiology, and psychophysics. Hear Res 242, 141-153.

[58] Srinivasan, A.G., Landsberger, D.M. & Shannon, R.V. 2010 Current focusing sharpens local peaks of excitation in cochlear implant stimulation. Hear Res 270, 89-100.

[59] Berndt, A., Yizhar, O., Gunaydin, L.A., Hegemann, P. & Deisseroth, K. 2009 Bi-stable neural state switches. Nat Neurosci 12, 229-234. (doi:10.1038/nn.2247).

[60] Kleinlogel, S., Feldbauer, K., Dempski, R.E., Fotis, H., Wood, P.G., Bamann, C. & Bamberg, E. 2011 Ultra light-sensitive and fast neuronal activation with the Ca(2)+-permeable channelrhodopsin CatCh. Nat Neurosci 14, 513-518. (doi:10.1038/nn.2776).

[61] van Hateren, J.H. 1990 Directional tuning curves, elementary movement detectors, and the estimation of the direction of visual movement. Vision Res 30, 603-614.

[62] Zeitz, C., Robson, A.G. & Audo, I. 2014 Congenital stationary night blindness: An analysis and update of genotype-phenotype correlations and pathogenic mechanisms. Prog Retin Eye Res. (doi:10.1016/j.preteyeres.2014.09.001).

[63] Cibis, G.W. & Fitzgerald, K.M. 2001 The negative ERG is not synonymous with nightblindness. Transactions of the American Ophthalmological Society 99, 171-176.

[64] Bijveld, M.M., Florijn, R.J., Bergen, A.A., van den Born, L.I., Kamermans, M., Prick, L., Riemslag, F.C., van Schooneveld, M.J., Kappers, A.M. & van Genderen, M.M. 2013 Genotype and phenotype of 101 dutch patients with congenital stationary night blindness. Ophthalmology 120, 2072-2081. (doi:10.1016/j.ophtha.2013.03.002).

[65] Robertson, D. & Irvine, D.R. 1989 Plasticity of frequency organization in auditory cortex of guinea pigs with partial unilateral deafness. J Comp Neurol 282, 456-471. (doi:10.1002/cne.902820311).

[66] Recanzone, G.H., Schreiner, C.E. & Merzenich, M.M. 1993 Plasticity in the frequency representation of primary auditory cortex following discrimination training in adult owl monkeys. J Neurosci 1993 Jan;13(1 87-103): 13, 87-103.

[67] Smirnakis, S.M., Brewer, A.A., Schmid, M.C., Tolias, A.S., Schuz, A., Augath, M., Inhoffen, W., Wandell, B.A. & Logothetis, N.K. 2005 Lack of long-term cortical reorganization after macaque retinal lesions. Nature 435, 300-307. (doi:10.1038/nature03495).

[68] Ghose, G.M., Yang, T. & Maunsell, J.H. 2002 Physiological correlates of perceptual learning in monkey V1 and V2. J Neurophysiol 87, 1867-1888.

[69] Baseler, H.A., Brewer, A.A., Sharpe, L.T., Morland, A.B., Jagle, H. & Wandell, B.A. 2002 Reorganization of human cortical maps caused by inherited photoreceptor abnormalities. Nat Neurosci 5, 364-370. (doi:10.1038/nn817).

[70] Levin, N., Dumoulin, S.O., Winawer, J., Dougherty, R.F. & Wandell, B.A. 2010 Cortical maps and white matter tracts following long period of visual deprivation and retinal image restoration. Neuron 65, 21-31. (doi:10.1016/j.neuron.2009.12.006).

[71] Fine, I., Wade, A.R., Brewer, A.A., May, M.G., Goodman, D.F., Boynton, G.M., Wandell, B.A. & MacLeod, D.I. 2003 Long-term deprivation affects visual perception and cortex. Nat Neurosci 6, 915-916. (doi:10.1038/nn1102).

[72] Masuda, Y., Dumoulin, S.O., Nakadomari, S. & Wandell, B.A. 2008 V1 projection zone signals in human macular degeneration depend on task, not stimulus. Cereb Cortex 18, 2483-2493. (doi:10.1093/cercor/bhm256).

[73] Baseler, H.A., Gouws, A., Haak, K.V., Racey, C., Crossland, M.D., Tufail, A., Rubin, G.S., Cornelissen, F.W. & Morland, A.B. 2011 Large-scale remapping of visual cortex is absent in adult humans with macular degeneration. Nat Neurosci 14, 649-655. (doi:10.1038/nn.2793).

[74] Baker, C.I., Peli, E., Knouf, N. & Kanwisher, N.G. 2005 Reorganization of visual processing in macular degeneration. J Neurosci 25, 614-618.

[75] Baker, C.I., Dilks, D.D., Peli, E. & Kanwisher, N. 2008 Reorganization of visual processing in macular degeneration: replication and clues about the role of foveal loss. Vision Res 48, 1910-1919. (doi:10.1016/j.visres.2008.05.020).

[76] Dilks, D.D., Baker, C.I., Peli, E. & Kanwisher, N. 2009 Reorganization of visual processing in macular degeneration is not specific to the "preferred retinal locus". J Neurosci 29, 2768-2773. (doi:10.1523/JNEUROSCI.5258-08.2009).

[77] Sunness, J.S., Liu, T. & Yantis, S. 2004 Retinotopic mapping of the visual cortex using functional magnetic resonance imaging in a patient with central scotomas from atrophic macular degeneration. Ophthalmology 111, 1595-1598. (doi:10.1016/j.ophtha.2003.12.050).

[78] Masuda, Y., Horiguchi, H., Dumoulin, S.O., Furuta, A., Miyauchi, S., Nakadomari, S. & Wandell, B.A. 2010 Task-dependent V1 responses in human retinitis pigmentosa. Invest Ophthalmol Vis Sci 51, 5356-5364. (doi:10.1167/iovs.09-4775).

[79] Georgeson, M.A. & Sullivan, G.D. 1975 Contrast constancy: deblurring in human vision by spatial frequency channels. J Physiol (Lond) 252, 627-656.

[80] Falconbridge, M., Wozny, D., Shams, L. & Engel, S.A. 2009 Adapting to altered image statistics using processed video. Vision Res 49, 1757-1764.

[81] Bao, M. & Engel, S.A. 2012 Distinct mechanism for long-term contrast adaptation. Proc Natl Acad Sci U S A 109, 5898-5903. (doi:10.1073/pnas.1113503109).

[82] Webster, M.A., Georgeson, M.A. & Webster, S.M. 2002 Neural adjustments to image blur. Nat Neurosci 5, 839-840. (doi:10.1038/nn906).

[83] Vaziri-Pashkam, M. & Cavanagh, P. 2008 Apparent speed increases at low luminance. J Vis 8, 1-12.

[84] Johnston, A. 2010 Modulation of time perception by visual adaptation. In Attention and Time (eds. A.C. Nobre & J.T. Coull), pp. 187-2000, Oxford University Press.

[85] Bruno, A., Ng, E. & Johnston, A. 2013 Motion-direction specificity for adaptation-induced duration compression depends on temporal frequency. J Vis 13, 19. (doi:10.1167/13.12.19).

[86] Ahuja, A.K., Behrend, M.R., Kuroda, M., Humayun, M.S. & Weiland, J.D. 2008 An in vitro model of a retinal prosthesis. Biomedical Engineering, IEEE Transactions on 55, 1744-1753.

# Figure Legends

Figure 1. A simulation of the perceptual information carried by ON- and OFF-centre pathways. (Column A) Two example images: a white-on-black eye-chart and a frame of a natural scene movie of a child scooting. Both subtend 24 degrees of visual angle. (Column B) These two images were convolved by a bank of difference of Gaussian filters followed by rectification to provide a rough approximation of the information carried by ON-centre pathways. (Column C) Images filtered by an inverted version of the same bank of Gaussian filters followed by rectification provides a rough approximation of the information carried by OFF-centre pathways. (Column D) The difference of the images of Panels B and C restore a band-pass version of the original image. (Column E) The sum of the images Panels B and C, represents an approximation of one potential perceptual outcome of simultaneously stimulating both ON-centre and OFF-centre pathways. Movie versions are shown for Panels A-E in Supplementary Materials.

Figure 2. An example of how a model of retinal axonal pathways can predict patient percepts. (A) Schematic of why axonal stimulation results in axon comets. Each image of the retina is flipped so that the upper region of the retina represents the upper visual field. The modelled axon trajectories (green lines) are based on a computational model of axon fibre trajectories developed using traced nerve fibre bundle trajectories extracted from fundus photographs of 55 human subjects. [51]. The red circles represent ganglion cell bodies whose axon fibres pass underneath the electrode on their way to the optic nerve, yellow shading represents a perceptual ‘axon-comet’. The dotted red box outlines the retinal regions shown in Panels B-D. (B, C) A subject implanted with the Argus 1 prosthesis (this array contained interleaved 250 and 500m electrodes) was simultaneously stimulated on two different pairs of electrodes. Each pair of stimulated electrodes are shown outlined in red. Predicted percepts generated using the model described above are shown in white. (D, E) Subject drawings (averaged over 5 trials) of the percepts induced by these stimulation patterns.

Figure 3. Simulations of perceptual distortions as a result of axonal stimulation. (A, B) Two images (the same as in Figure 1) are overlaid on the retinal surface. The position of the electrode array (subtending 12 degrees) is shown by a red dotted box. (C, D) The predicted effect of axonal stimulation on the two images of Figure 1B and D for λ = 0.5, 1 and 2. Movie versions are shown in Supplementary Materials.

Figure 4. Example of one test within a standardized ‘Activities of daily living’ test proposed by [54, reprinted with their kind permission]. The subject is asked to identify, describe and localize objects while sitting at a table.

Figure 5. The perceptual effects of sluggish response kinetics. (A) Response kinetics for LiGluR-MAG0460 and normal phototransduction. LiGluR-MAG0460 cells are activated by visible light (time-course shown as a yellow outlined bar) and relax spontaneously in the dark. Blue circles show human embryonic kidney cell recordings in voltage-clamp configuration at −75 mV [replotted with kind permission from 7]. Our simulated approximation (blue solid line) used an exponential onset of time-constant of 20ms and an offset time-constant of 200ms. For comparison, a crude simulation of primate phototransduction dynamics (using an exponential time-constant of 20ms for both onset and offset) is also plotted (red line). For our purposes a crude approximation of time-courses was adequate, more accurate modelling of these time-courses can be found in [7] and [61] respectively. (B) The image of a child scooting (same as Figure 1A). (C) The image of Figure 1A was in fact taken from a movie sequence. Here we show the image frame of 1A, with the movie filtered using the temporal dynamics of LiGluR-MAG0460. Rather dramatically, the sluggish temporal dynamics causes the scooting child to almost completely disappear. The movie version of Panel C is shown in Supplementary Materials.

Figure 6. Three example models of the potential perceptual experience of sight recovery. All images subtend 12 degrees of visual angle. (A) Scoreboard model. The position of the array is again shown with red dotted lines. The luminance of the apparent percept is linearly related to the strength of current on the retina. (B) Simulation of electrical stimulation. This particular simulation is based on the model of simultaneously stimulating ON- and OFF pathways as described in Figure 1E, followed by the model of axonal stimulation as described in Figure 3 (C) Simulation of small molecule photoswitch stimulation. This simulation is based on the model of simulating ON-centre pathways in isolation as described in Figure 1B, followed by the effects of sluggish temporal dynamics as described in Figure 5. Movie versions are shown in Supplementary Materials.

# Supplementary Movie Legends

Movie Examples for Figure 1. Simulations of the perceptual information carried by ON- and OFF-centre pathways. (A) A natural scene movie of a child scooting subtending 24 degrees of visual angle is used as an example. (B) Each frame was convolved by a bank of difference of Gaussian filters followed by rectification to provide a rough approximation of the information carried by ON-centre pathways. (C) Each frame was filtered by an inverted version of the same bank of Gaussian filters followed by rectification to provide a rough approximation of the information carried by OFF-centre pathways. (D) The difference of the images of Panels B and C restore a band-pass version of the original image. (E) The sum of the images Panels B and C, represents an approximation of one potential perceptual outcome of simultaneously stimulating both ON-centre and OFF-centre pathways.

Movie Examples for Figure 2. Simulations of perceptual distortions as a result of axonal stimulation. The predicted effect of axonal stimulation are shown for (A) λ = 0.5, (B) λ = 1 (C) λ = 2.

Movie Example for Figure 5. The perceptual effects of sluggish response kinetics using the temporal dynamics of LiGluR-MAG0460.

Movie Examples for Figure 6. Three example models of the potential perceptual experience of sight recovery. All images subtend 12 degrees of visual angle. (A) Scoreboard model. The luminance of the apparent percept is linearly related to the strength of current on the retina. (B) Simulation of electrical stimulation. This particular movie is based on the model of simultaneously stimulating ON- and OFF pathways as described in Figure 1E, followed by the model of axonal stimulation as described in Figure 3. (C) Simulation of small molecule photoswitch stimulation. This movie is based on the model of simulating ON-centre pathways in isolation as described in Figure 1B, followed by the effects of sluggish temporal dynamics as described in Figure 5.